## Sheet1
SOH OPD & IPD MEDICAL EQUIPMENT TAGGING
Tag No. | Equipment Description | Additional Comments | Brand name | Electrical Specifications | Procurement/ Delivery Schedule | Installation/ Commisioning Schedule | TANIT Comments | Location
4 | EXAMINATION HALOGEN LIGHT - ON CASTORS |  |  | Equipment is mobile on wheels, Power supply: 220V, 50Hz, 20W | 2023-01-20 | 2023-02-01 | OK | OPD
5 | DIAGNOSTIC STATION - WALL MOUNTED - OPHTHALMOSCOPE - OTOSCOPE - NIBP - THERMOMETER |  | Welch Allen | Equipment is wall mounted; Electrical specifications: 100 – 240 V AC, 50 – 60 Hz, 1.5 – 0.8 A | 2023-01-20 | 2023-02-01 | OK | OPD
6 | ULTRA SLIM LIGHT BOX 41X36 CM |  | GF TENDERCARE | Equipment is Table top. Input Power Supply: 100-230V, 50-60Hz, 120VA | 2023-01-20 | 2023-02-01 | OK | OPD
7 | ECG - 12 CHANNEL - ON TROLLEY |  | GE Mac 2000 | Equipment is mounted on cart to a height of 200mm. Equipment uses AC/DC or Battery Operation. Input Power Supply: 100-240VAC, 47-63Hz, maximum 1.5A. Equipment is also LAN enabled. | 2023-01-20 | 2023-02-01 | OK | OPD
8 | DEFIBRILLATOR - UNMONITORED |  | Progetti | Equipment is mobile, uses rechargeable battery. AC charger power supply, Input: 100 ~ 240V AC 50/60Hz max. Battery pack: 14.4V – 2Ah Ni-Mh battery (internal rechargeable) charging time maximum 3 hours | 2023-01-20 | 2023-02-01 | OK | OPD
9 | TREADMILL ELEVATING WITH PC CONTROL |  | GE Case + T2100 Treadmill | 200 to 240 VAC, 50/60 Hz, single-phase, 20 amperes, NEMA 6-20 R wall socket (dedicated circuit recommended | 2023-01-20 | 2023-02-01 | OK | OPD
10 | ULTRASOUND UNIT - OB/GYN 3/4D VOLUMETRIC WITH DVD RECORDER |  | Voluson S10 | 13A Socket outlet wired with >3 core 2.5mm2 cable, Power Consumption: 900VA Earthing system of less than 2ohms, Voltage 220-240VAC. single phase UPS for Ultrasound; Dedicated AC Power outlet must be within reach of the system without extension cords within 1m of the system. Airconditional Availability | 2023-01-01 | 2023-01-01 | OK | OPD
11 | PLETHYSMOGRAPH CABINET - COMPACT SIZE WITH CHAIR |  | nSpire | Input Power Supply: 100-240V ± 10%, 50/60 Hz. Box Dimension: 93*85*172cm Equipment requires external PC connection. | 2023-01-20 | 2023-02-01 | OK | OPD
12 | ELECTROENCEPHALOGRAPH STATION - 32 CHANNEL TROLLEY MOUNTED - WITH ACCESSORIES |  | Neurowork | Power frequency (50/60 Hz) magnetic field IEC 61000-4-8 | 2023-01-01 | 2023-02-01 | OK | OPD
13 | ENT STATION WITH CHAIR AND INSTRUMENTS - VIDEO MONITOR | Total station inculding Chair | Euroclinic | 220/240 V(110/115 VAC) | 2023-01-01 | 2023-02-01 | OK | OPD
14 | ENT MICROSCOPE |  | Leica | Equipment is mobile on caster wheels. Power supply: 100-240V AC, 50/60Hz. Power consumption: 1200VA | 2023-01-01 | 2023-02-01 | OK | OPD
15 | ELECTROSURGICAL UNIT 200 WATTS - ENT | On ENT Station | Inspital | Input Power Voltage: 100-240VAC, 50-60Hz Input Power: 750VA Maximum Output Power 200W | 2023-01-01 | 2023-02-01 | OK | OPD
18.2 | SCREENER - VISUAL FIELD COMPUTER BASED |  | LIGHT BOX | Equipment is Table top, LAN Connectivity enabled. Input Power Supply: 85 – 263 V AC, 50/60 Hz, 60 VA | 2023-01-01 | 2023-02-01 | OK | OPD
19 | CEPHALO SCAN CBCT |  | Belmont | Power Requirement x-ray will operate on a power supply of 120 + 10 VAC. A three wire GROUNDED circuit, separately connected to the central distribution panel with an over current protection device rate for 20 amperes. Recommended wire size is 12 AWG. But if the wire run distance is to exceed 50 feet, 10 AWG is required. For wire run distance in excess of 75 feet, upto 125 feet, 8 AWG is required. Line voltage regulation must exceed 3 % at 12.5 amps. | 2023-01-01 | 2023-02-01 | OK | OPD
20 | DENTAL CHAIR BUILT IN COMPRESSOR COMPLETE |  | Olsen | Equipment is powered by a Motor: 24V, 46W, 45rpm | 2023-01-01 | 2023-02-01 | We need shop drawings or more details about drain location, water inlt location, etc. to prepare the site. | OPD
21 | DENTAL X-RAY UNIT WALL MOUNTED |  | Belmont | Maximum Length Reach of equipment : 84-1/8” with 40.0” arm Power Supply Electrical requirement Power supply of 120VAC ± 10% (108 ~ 132V AC) with a 3 wire GROUNDED circuit, separately connected to the central distribution panel with an over current protection device rated for 15 amperes. Recommended wire size is 14 AWG ; but if the wire run distance is to exceed 50 feet (15m), 12 AWG is required. For wire run distances in excess of 75 feet (23m) 10 AWG is required. Line voltage regulation must be within 0 ~ 3 % at 6.5 amps. Please check drawing on separate sheet | 2023-01-01 | 2023-01-01 | OK | OPD
22 | DENTAL CURING LIGHT - CORDLESS - LED |  | Woodpecker | Equipment is mobile, uses Rechargeable lithium battery. Power input: 100V to 240V, 50/60Hz, 10VA | 2023-01-01 | 2023-02-01 | OK | OPD
23 | AUTOCLAVE - DENTAL INSTRUMENTS 24L |  | Tuttnauer/ 2840EKA D(28L) | Voltage/ Frequency (Hz): 230V/60Hz Power: 1400W 6A Equipment is Table Top with Dimension 630*530*440mm | 2023-01-01 | 2023-02-01 | OK | OPD
24 | LASER - DENTAL SURGICAL/DIODE/TABLETOP FOR SOFT TISSUE |  | Zolar | Equipment is table-top, uses Battery AC100-240V 50/60Hz Wireless Pedal 1.5V Battery | 2023-01-01 | 2023-02-01 | OK | OPD
25 | AMALGAMATOR |  | Henry Schein | Equipment is Table-top Input Voltage: 110V, 60Hz Power Output: 40W | 2023-01-01 | 2023-02-01 | OK | OPD
26 | RETINAL CAMERA / OCT OPHTHALMOSCOPE / EYE F-ANGIO / TABLETOP |  | Topcon | Equipment is Table-top Input Voltage: 110=240V, 50-60Hz Power Consumption: 400VA (maximum) 100VA (normal) Dimensions W*D*H: (274*508*566mm) | 2023-01-01 | 2023-02-01 | OK | OPD
27 | PROJECTOR - VISUAL ACUITY |  | LIGHT BOX AUTOMATED ACUITY CHAT | Equipment is Table top, can also be mounted on wall. Input Power Supply: 115-230 VAC 50/60 Hz | 2023-01-01 | 2023-02-01 | OK | OPD
28 | AUTO REFRACTOMETER/KERATOMETER |  | REICHART | Equipment is Table top. Input Power Supply: 100-240V, 50/60Hz, 60VA | 2023-01-01 | 2023-02-01 | OK | OPD
28.1 | LENSOMETER - TABLETOP |  | LIGHT BOX | Equipment is Table top. Input Power Supply: 6VAC Adapter 220V | 2023-01-01 | 2023-02-01 | OK | OPD
29 | OPHTHALMOLOGY CHAIR - COMPLETE (RETINOSCOPE - SLIT LAMP - TONOMETER) |  | Light Box | Input Power Supply: 110-230V, 50/60Hz | 2023-01-01 | 2023-02-01 | OK | OPD
30 | LED VIDEO COLPOSCOPE - 3.75X/7X/15X > 28000 LUX - ON TROLLEY |  | Ecleris/ C100-FID Colposcope | Equipment is mobile on caster wheels. Power supply: 100-240V AC, 50/60Hz. | 2023-01-01 | 2023-02-01 | OK | OPD
31 | BOOTH - AUDIOMETRY - COMPACT WITH VIEWPORT AND CHAIR - 1.2X1.2M |  | sibelmed | Input Power Supply: 220 V 50-60 Hz, 50W | 2023-01-01 | 2023-02-01 | OK | OPD
32 | ERGOMETER |  | GE | Power Supply: 100 to 240 V, 50 to 60 Hz Power Consumption: 60VA max | 2023-01-01 | 2023-02-01 | OK | OPD
33 | ELECTROTHERAPY UNIT |  | Doctor life | Input Power Supply: 220 V 50-60 Hz, 50W | 2023-01-01 | 2023-02-01 | OK | OPD
33.1 | PRESSOTHERAPY UNIT |  | EME Physio | Input Power Supply: 220 V 50-60 Hz, 50W | 2023-01-01 | 2023-02-01 | OK | OPD
33.2 | VACCUM THERAPY |  | Motech | Equipment is mobile, hand-held Input Power Supply: 200 V, 50Hz | 2023-01-01 | 2023-02-01 | OK | OPD
33.3 | INFRARED RADIANT |  | Motech | Equipment is mobile, hand-held Input Power Supply: 200 V, 50Hz | 2023-01-01 | 2023-02-01 | OK | OPD
33.4 | MAGNETOTHERAPY |  | Chinesport | Equipment is Table-top Input Power Supply: 230 VAC, 50-60 Hz | 2023-01-01 | 2023-02-01 | OK | OPD
33.5 | CONTINUOUS PASSIVE MOTION UPPER LIMBS |  |  | Power supply: 220V AC, 50Hz. Input Power: <60VA | 2023-01-01 | 2023-02-01 | OK | OPD
33.6 | CONTINUOUS PASSIVE MOTION LOWER LIMBS |  | EME Physio | Input Power Supply: 220V, 50Hz | 2023-01-01 | 2023-02-01 | OK | OPD
33.7 | HYDROCOLLECTOR - COLD COMPRESS |  |  | Power Supply: 100 to 240 V, 50 to 60 Hz Power Consumption: 1500W max Dimension: 67*40*84cm | 2023-01-01 | 2023-02-01 | OK | OPD
33.8 | HYDROCOLLECTOR - HOT COMPRESS |  |  |  | 2023-01-01 | 2023-02-01 | OK | OPD
34 | ULTRASOUND UNIT - CARDIAC - WITH ASSESSMENT SOFTWARE - 3 PROBE |  | GE Vivid S7 | 13A Socket outlet wired with >3 core 2.5mm2 cable, Power Consumption: 0.5-1A Earthing system of less than 2ohms, Voltage 240 VAC single phase UPS for Ultrasound Airconditional Availability | 2023-01-01 | 2023-01-01 | OK | OPD
35 | COMPUTER-BASED URODYNAMIC SYSTEM ON CASTORS |  | creomidical | Equipment is moveable on castor. LAN Connection available. Input Power Supply: 220V, 50Hz | 2023-01-01 | 2023-01-01 | OK | OPD
36 | URODYNAMIC CHAIR |  | Creo Medical | Input Power Supply: 220-240V, 50/60Hz | 2023-01-01 | 2023-01-01 | OK | OPD
37 | REFRIGERATOR 500 L |  | Haier | Equipment is mobile on caster wheels. Power supply: 100-240V AC, 50/60Hz. | 2023-01-01 | 2023-01-01 | OK | OPD
38 | PORCELAIN FURNACE WITH VACCUM PUMP |  | Phoenix 7 | Input Power Supply: 220V, 1.6W | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
39 | BURNOUT FURNACE |  | UMG | Input Power Supply: 240V, 13A, 3.1kW, 1 Ph, 50/60Hz; Power Consumption: 4000W | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
40 | CASTING MACHINE - MANUAL |  | Mancast | Equipment is table top, non electrical. | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
41 | SANDBLASTER |  | Aixin | Input Power Supply: 220V, 50/60Hz Power Consumption: <8W | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
42 | MICRO SANDPLASTER |  | Optiblast | Equipment is table top, mobile. Input Power Supply: 220V, 50Hz | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
43 | VIBRATOR |  | European | Equipment is table top Input Power Supply: 220V, 50Hz | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
44 | HANGING MOTOR WITH HANPIECE |  | European | Equipment is mobile with rechargeable battery 40V, 2.5Ah Input Power Supply: 220V, | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
45 | MICRO MOTOR |  | European | Equipment is attached to motor | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
46 | VACCUM MIXER |  | China | Equipment is table top Input Power Supply: 110/230V, 50/60Hz. Motor: 320RPM | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
47 | PINDEX |  | China | Equipment is table top Input Power Supply: 230V, 50Hz, 2500rpms Nominal Power: 100W | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
48 | MODEL TRIMMER |  | European | Power Supply: 220V/110V 50/60Hz Power Consumption: 370W Rotational Speed: 2,800RPM | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
49 | STEAM MACHINE |  | China | Power Supply: 120/220V Power Consumption: 10amps | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
50 | MOTOR CENTRIFUGE |  | UK | Power Supply: 220V | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
51 | POLISHING MOTOR |  | European | Power Supply: 220V, 50/60Hz | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
52 | INSTRUMENT CLEANER |  | MULIT ENZYMATIC | Consumable, Non-Electrical Item | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
53 | SEALING MACHINE |  | QASA | Equipment is Table top Power Input: max 420W | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
54 | ORTHODONTIC SPOT WELDER |  | Vimel | Equipment is Table top Power Input Voltage: 220V, 50/60Hz Supply current on load: 15A Power: 1000W | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
55 | VACUUM FORMER TABLETOP DENTAL LAB |  | UMG | Equipment is Table top Power Input Voltage: 220V, 50/60Hz Power: 350W | 2023-01-01 | 2023-01-01 | Confirm equipment positioning shared by TANIT earlier is going to be followed | OPD
56 | ELECTRONIC TRANSPORT VENTILATOR ON BED HOOK |  | Siare | Power Requirement: 100-240 VAC, 47-63 Hz Power consumption: 50 Watts Internal battery | 2023-03-01 | 2023-03-01 | OK | IPD | 50
57 | NEBULISER - EMERGENCY ROOM WITH MEDICATION CUP | table Top | Omron | Equipment is Table top Power Requirement: 100-240V ~ 350mA, 50/60 Hz | 2023-03-01 | 2023-03-01 | OK | IPD | 250
58 | PATIENT WARMER |  | Nellcor | Equipment is mobile on castor. Heater Size: 1000W Power Requirements 240 V AC Max. current at 240 V = 5 A Input Frequency 50/60 Hz ±1 Hz Fuse (x2) Littlefuse (mfr.) 0218010; 250 V, 10 A | 2023-03-01 | 2023-03-01 | OK | IPD | 1000
59 | VITALS MONITOR ON CASTORS - DETACHABLE |  | Mindray | Power Requirements: 100 to 240 VAC, 0.9 to 0.5 A, 50/60 Hz. Battery Type: Rechargeable lithium ion. Run Time: Up to 8 hours when powered by a new fully-charged battery at 77° F ± 41° F (25° C ± 5° C) with SpO2 and NIBP measurements at an interval of 15 minutes. | 2023-03-01 | 2023-03-01 | OK | IPD
60 | TRANSPORT & EMERGENCY MONITOR |  | GE | Power Requirement: 100-240 VAC, 50-60 Hz Power consumption: 300 Watts(Max) | 2023-03-01 | 2023-03-01 | OK | IPD
61 | DEFIBRILLATOR - MONITORED WITH PACER |  | PROGETTI | Equipment is mobile AC Power Supply input: 240V, 60Hz Also uses rechargeable internal lithium-ion battery | 2023-03-01 | 2023-03-01 | OK | IPD
62 | PUMP - SUCTION ON TROLLEY - TWIN CONTAINER 1L |  | Yuwell | Equipment is moveable on castor Power Input Voltage: AC220V +22V, 50Hz +1Hz | 2023-03-01 | 2023-03-01 | OK | IPD
63 | AUTO SYRINGE PUMP |  | PROGETTI | Power Supply AC power supply: 100-240V 50/60Hz Input power: 50VA DC power supply: DC 15V lithium battery | 2023-03-01 | 2023-03-01 | OK | IPD
64 | INFUSION PUMP VOLUMETRIC ON CASTORS |  | PROGETTI | Power Supply: AC Power supply: AC 240V, 60Hz Input Power: 50VA DC power supply: DC 15V Lithium Battery | 2023-03-01 | 2023-03-01 | OK | IPD
65 | SAW - PLASTER – ELECTRIC |  | AESCULAP | Power Supply: AC Power supply: AC 240V, 50Hz | 2023-03-01 | 2023-03-01 | You shared a full catalogue. Which model are you providing? | IPD
66 | SAW - PLASTER – MANUAL |  | STATRIC | Equipment is non-electrical hand held plaster saw | 2023-03-01 | 2023-03-01 | OK | IPD
67 | AUTOCLAVE - FAST CYCLE - 30L - HORIZONTAL - TABLE TOP |  | BIOBASE | Power Supply Voltage: 110/220V 50/60Hz Power Consumption: 1700W Equipment is table top | 2023-03-01 | 2023-03-01 | Required 30L. Which model are you providing? | IPD
68 | SLOP HOPPER |  | XINHE/CHBC | Equipment is non electrical. Needs water inlet and drain | 2023-03-01 | 2023-03-01 | We need more details: exact specifications and locations of water inlet and drain | IPD
69 | ANAESTHESIA UNIT - 2 VAPORIZER/VENTILATOR/RESP MONITOR |  | GE | Power Supply: AC Power supply: AC 240V, 50Hz Equipment uses storage battery. | 2023-03-01 | 2023-03-01 | OK | IPD
70 | TABLE - OPERATING - MINOR SURGERY - HYDRAULIC |  | PROGETTI | Equipment is non electrical | 2023-03-01 | 2023-03-01 | How is it non electrical and both models on provided catalogue (PG OT1-ONE and PG OT1-ONE Plus) have a remote control and a power consumption? | IPD
71 | CEILING LIGHT - SURGICAL - LED - 100K LUX - SINGLE ARM |  | INSPITAL | Equipment is ceiling mount with back-up battery | 2023-03-01 | 2023-03-01 | Which model are you providing? What are the technical specs? | IPD
72 | WASHER - BEDPAN/URINAL - FLOOR TYPE |  | STANDBRIDGE | Non electrical equipment | 2023-03-01 | 2023-03-01 | We need additional details that show it is non electrical. No Datasheet provided | IPD
73 | GENERAL ULTRASOUND UNIT 3 PROBE WITH PRINTER |  | GE Logic P9 | Electrical power Voltage: 100 – 240 VAC Frequency: 50/60 Hz Power consumption: No more than 200 VA with peripherals | 2023-03-01 | 2023-03-01 | OK | IPD
74 | PRINTER NETWORK BASED LASER B/W |  | HP | Equipment is table top Power Supply Voltage: 220V, 60Hz, 2.7A Power Consumption: 420W (printing), 2.0W (Ready) | 2023-03-01 | 2023-03-01 | A bigger printer (Floor Standing) is requested with scan capabilities | IPD
75 | FILM PRINTER MULTI FORMAT FOR RADIOLOGY DEPARTMENT |  | Fuji | Power Supply Conditions: Single phase 50-60Hz 120-240V ±10% 5A (max) | 2023-03-01 | 2023-03-01 | OK | IPD
76 | REPORTING STATION (WITH THE DESIRED 3-WORKSTATIONS ACCORDING TO STATED SPECIIFCATION) |  | GE Workstation | Electrical power Voltage: 100 – 240 VAC Frequency: 50/60 Hz | 2023-03-01 | 2023-03-01 | OK | IPD
77 | ULTRA SLIM LIGHT BOX DOUBLE FILM |  | TEMED | Power Supply Conditions: 110-230V, 50/60Hz | 2023-03-01 | 2023-03-01 | Does it need a cable or a power socket? And at which height? | IPD
78 | BED 3 SECTION ELECTRIC - FIXED HEIGHT - ON CASTORS |  | IMO | Equipment is moveable on wheels with back-up battery Power Supply: 110-240V, 50/60Hz Control Unit Power: 390VA Battery Charge: 18W | 2023-03-01 | 2023-03-01 | Please confirm it comes with a mattress. | IPD
79 | BED INTENSIVE CARE FULL ELECTRIC |  | IMO | Equipment is moveable on wheels with back-up battery Power Supply: 110-240V, 50/60Hz Control Unit Power: 390VA Battery Charge: 18W | 2023-03-01 | 2023-03-01 | Matrix E20 is rejected. Higher specs for ICU are needed. Mattress to be included | IPD
80 | MONITOR - PATIENT ICU - ECG/SPO2/NIBP/RESP/ETCO2/TEMP |  | GE | Power Requirement: 100-240 VAC, 50-60 Hz Power consumption: 300 Watts(Max) | 2023-03-01 | 2023-03-01 | OK | IPD
81 | VENTILATOR - ADULT - ICU |  | Siare | Power Supply: 100-240VAC/ 47-63Hz Power: 50W Internal battery: 2 Batteries (Pb 12VDC, 1,3Ah) | 2023-03-01 | 2023-03-01 | OK | IPD
82 | NEBULISER - EMITTED DOSE - ICU/CCU |  | Omron | Equipment is mobile Rating (AC adapter): 100 - 240V ~ 350mA, 50/60 Hz Rating (Compressor Nebulizer): 12V 0.8A | 2023-03-01 | 2023-03-01 | OK | IPD
83 | VENTILATOR - BIPAP CAPABILTY - ICU |  | Siare | Electric power supply 100 ÷ 240Vac / 47 ÷ 63Hz Power Max 150 Watt External power supply (low tension) 12 Vdc / 7 A Internal battery Battery NiMh 12Vcc – 4.5 Ah | 2023-03-01 | 2023-03-01 | OK | IPD
84 | VENTILATOR - TRANSPORT - ADULT |  | Siare | Electric power supply 100 ÷ 240Vac / 47 ÷ 63Hz Power Max 150 Watt External power supply (low tension) 12 Vdc / 7 A Internal battery Battery NiMh 12Vcc – 4.5 Ah | 2023-03-01 | 2023-03-01 | OK | IPD
85 | ANALYZER - BLOOD GAS WITH ELECTROLYTES STAT FOR ICU |  | Istat | Equipment is handheld | 2023-03-01 | 2023-03-01 | OK | IPD
86 | ULTRASOUND UNIT - PORTABLE - SINGLE PROBE |  | GE Versana Active | Electrical power Voltage: 100 – 240 VAC Frequency: 50/60 Hz Power consumption: No more than 200 VA with peripherals | 2023-03-01 | 2023-03-01 | OK | IPD
87 | X-RAY MACHINE - MOBILE - ADULT |  | GE Optima 240AMX | Equipment is mobile, LAN connection enabled. Operates with internal battery. Power Capacity: 100-240 V nominal, 50/60 Hz operation | 2023-03-01 | 2023-03-01 | OK | IPD
88 | DIALYSIS MACHINE - ICU |  | GAMBRO PRISMAFLEX | Supply 100-240 +/– 10% Vac, 50/60 Hz Maximum power consumption 150W Normal average power consumption 150 W Input line current: 5A max rms at 100 Vac, 2.5A max rms at 240 Vac | 2023-03-01 | 2023-03-01 | OK | IPD
89 | MONITOR - PEDIATRIC ICU |  | GE | Electrical Specification: AC Input: 100 to 240V, 50/60Hz, 150VA Back up battery: Exchaneable Lithium-ion 2 pcs | 2023-03-01 | 2023-03-01 | OK | IPD
90 | SCRUB STATION-TRIPLE WIDTH-SENSOR OPERATED-SOAP DISPENSER |  | INSPITAL | Equipment is non electrical | 2023-03-01 | 2023-03-01 | Which model are you providing? Why is it non electrical? It should be sensor operated | IPD
91 | NEBULISER | Table Top | Omron | Rating (AC Adapter): 100-240V , 50/60Hz, 0.12-0.065A | 2023-03-01 | 2023-03-01 | OK | IPD
92 | ANAESTHESIA WORKSTATION- 2 VAPORIZER/DIGITAL/RESP+PATIENT MONITOR/VENTILATOR/GAS SCAVENGING | Equipment is mobile on castor | GE | Current leakage 100/120 V: < 500μA 220/240 V: < 500μA Power and battery backup Supply voltage: 100-120 Vac, 50/60 Hz | 2023-03-01 | 2023-03-01 | Are you providing Aisys or Carestation 30?? | IPD
93 | TABLE - OPERATING GENERAL SURGERY ELECTRIC/HYDRAULIC WITH MANUAL OVERRIDE - KIDNEY BRIDGE |  | PROGETTI | Power Supply: 100-240V , 50/60Hz, with Back-up battery | 2023-03-01 | 2023-03-01 | OK | IPD
94 | LIGHT - CEILING MAJOR SURGERY 2 ARMS 300 KLUX |  | INSPITAL |  | 2023-03-01 | 2023-03-01 | Only Catalogue Provided. No technical Specs. We need electrical details. | IPD
95 | PENDANT - CEILING / 2 TV |  | INSPITAL | Power Supply: 100-240V , 50/60Hz | 2023-03-01 | 2023-03-01 | This should be mounted on one of the surgical light arms | IPD
96 | DRILL - ELECTRICAL - SURGICAL - ORTHOPAEDIC - SET OF ATTACHMENTS - ASSORTED | handheld equipment | Ruijin | Equipment is hand-held with in-built battery. Working Voltage: 14.4V | 2023-03-01 | 2023-03-01 | OK | IPD
97 | DRILL - ELECTRICAL - SURGICAL - CRANIAL PLASTY - WITH CRANIAL PERFORATOR |  | RUIJIN | Charger input power AC 220V, 50Hz Charger output power 14.4V, 700mA Charger input power 30va Battery voltage 14.4V Battery charging time about 3 hours | 2023-03-01 | 2023-03-01 | OK | IPD
98 | PUMP - SUCTION ON TROLLEY TWIN CONTAINER 3L ELECTRIC |  | Yuwell | Equipment is moveable on castor Power Input Voltage: AC220V +22V, 50Hz +1Hz | 2023-03-01 | 2023-03-01 | OK | IPD
99 | MONITOR - PATIENT - OPERATING THEATER - ANA GASES/FULL PARAMETER | Mobile Equipment | GE | Power Input: 100 to 240V ±10%, 50/60 Hz | 2023-03-01 | 2023-03-01 | OK | IPD
100 | TOWER - ARTHROSCOPY COMPLETE - VIDEO PROCESSOR AND INSTRUMENT SETS |  | Mediteknik | Equipment is mobile on castor | 2023-03-01 | 2023-03-01 | Equipment needs power. Please provide Specifications. Catalogue is not enough | IPD
101 | C-ARM 12" HF FLUOROSCOPY - TRI FOCUS - 8KW |  | GE Brivo 785 | Electrical • Input power (60 Hz or 50 Hz) - 100V/110V/120V @ 20A - 200V @ 12A - 220V/230V/240V @ 10A | 2023-04-01 | 2023-04-01 | OK | IPD
102 | ELECTROSURGICAL UNIT 400 WATTS - DOUBLE SURGEON - VESSEL SEALING |  | INSPITAL | Maximum Power Output- 400 W-300 Ω | 2023-03-01 | 2023-03-01 | OK | IPD
103 | TOWER - LAPARASCOPY COMPLETE MONITOR/CAMERA/LIGHT SOURCE/INSUFFLATOR/PUMP/VIDEO PROC |  | Mediteknik | Equipment is mobile on castor Power Supply: 100-240V Power Consumption: 145VA | 2023-03-01 | 2023-03-01 | OK | IPD
104 | CRYOSURGERY UNIT OPHTHALMOLOGY - SET OF ACCESSORIES |  | Retikare | Electrical power: AC 100-240 Volts , 50/60 Hz, Single Phase | 2023-03-01 | 2023-03-01 | OK | IPD
105 | MICROSCOPE - OPERATING - NEUROSURGERY |  | leica | Power Connection: 1600 VA 50/60HZ 100 V, 120 V, 220 V, 240 V (+10 %/–15 %) 2 × T10 AL 100/120 V 2 × T8 AL 220/240 V | 2023-03-01 | 2023-03-01 | OK | IPD
106 | HYPO HYPERTHERMIA UNIT |  | Baer Hugger | Power Rating: 220-240 VAC, 50/60 Hz, 7.2 Amperes 100 VAC, 50/60 Hz, 14.5 Amperes | 2023-03-01 | 2023-03-01 | OK | IPD
107 | OPHTHALMOSCOPE - INDIRECT |  | Neitz | Equipment is mobile Voltage Input: AC120/230V | 2023-03-01 | 2023-03-01 | Are you providing Neitz or Welch Allen? | IPD
108 | PHACO EMULSIFIER UNIT | table top | MTP2000 | Equipment is Table top Power Rating: 100-240 VAC, 50/60 Hz | 2023-03-01 | 2023-03-01 | OK | IPD
109 | TROLLEY WASHER | Attached to IV Pole | CataRhex3 | Equipment is attached to IV pole Power Rating: 100-240 VAC | 2023-03-01 | 2023-03-01 | A Trolley washer is a stand alone equipment for washing trolleys in the CSSD. You provided a phaco catalogue for this. | IPD
110 | WASH STATION - DOUBLE SINK - STAINLESS STEEL - SPRAY GUN WATER | Wall mounted, Non electrical | Inspital | Equipment is wall mounted, non-electrical | 2023-03-01 | 2023-03-01 | Which model are you providing? We need the full details including shop drawings (Water inlet and drain for wash station as well as for the spray gun) | IPD
111 | WASHER - DISINFECTOR - AUTOMATIC CYCLES - PASS THROUGH |  | Biobase | Power Supply: 380V, 50Hz (Stanard), 110V, 60Hz (Optional) Power Consumption: 17KW | 2023-03-01 | 2023-03-01 | This is a glassware washer. We need CSSD washers disinfectors, pass through (Double doors) | IPD
112 | CLEANER - ULTRASONIC - >30L |  | Biobase | Power Supply: AC 220V, 50/60Hz | 2023-03-01 | 2023-03-01 | OK | IPD
113 | SPRAY GUN - AIR |  | UGM | Hand held equipment | 2023-03-01 | 2023-03-01 | We need additional details. Shop drawings that show the compressed air requested and the height. | IPD
114 | WORK TABLE - STAINLESS STEEL - 1.8 M - PACKING AND SORTING | Non electrical | Inspital | Hyraulic wark table, non electrical | 2023-03-01 | 2023-03-01 | This is a Mayo Table for Operating Rooms. We need a packing Table for CSSD !! | IPD
115 | EXAMINATION LIGHT - CSSD - MAGNIFYING - DESKTOP |  | INSPITAL | Power Supply: AC 110/220V, 50/60Hz | 2023-03-01 | 2023-03-01 | This is a surgical light. We need a magnifying desktop light for CSSD !! | IPD
116 | THERMOWELDER - HEAT SEALING - MANUAL | table top Equipment | Biobase | Power Supply: AC 110/220V, 50/60Hz Power Consumption: 200W | 2023-03-01 | 2023-03-01 | OK | IPD
117 | AUTOCLAVE - 500 L - DOUBLE DOOR - WITH LOADING AND UNLOADING TROLLEYS |  | BIOBASE | Power Supply: AC 220V, 50/60Hz Power Consumption: 7KW | 2023-03-01 | 2023-03-01 | Which one are you providing? We need a double door CSSD steam sterilizer | IPD
118 | CEILING LIGHT - SURGICAL - LED - 100K LUX - SINGLE ARM | Ceiling Mounted | INSPITAL | Equipment is ceiling mount with back-up battery | 2023-03-01 | 2023-03-01 | Which model are you providing? It should be electrical sinice it is ceiling mounted. | IPD
119 | EXAMINATION LED LIGHT - WALL MOUNTED |  | PROGETTI | Power Supply: 220V, 50Hz / 120V, 60Hz LED Power: 108W | 2023-03-01 | 2023-03-01 | Are you providing Progetti or Daray? | IPD
120 | SCRUB STATION-DOUBLE WIDTH-SENSOR OPERATED-SOAP DISPENSER | Wall Mounted | INSPITAL | Equipment is wall mounted, Power Supply: 100-240V | 2023-03-01 | 2023-03-01 | Which model?? We need shop drawings to prepare for the outlets | IPD
121 | ULTRASOUND UNIT - OB/GYN 3 PROBE WITH ENDOCAVITY AND BIOPSY SET |  | GE | Electrical power Voltage: 100 – 240 VAC Frequency: 50/60 Hz Power consumption: No more than 200 VA with peripherals | 2023-03-01 | 2023-03-01 | OK | IPD
122 | PATIENT CHAIR - RECLINED - 3 SECTION - HYDRAULIC - FIXED HEIGHT - ARM SUPPORT |  | Diamed | Mobile Hydraulic Chair, Not Electrical | 2023-03-01 | 2023-03-01 | OK | IPD
123 | DIALYSIS MACHINE |  | AK 98 (BAXTER) | Equipment uses a Water Treatment System (WRO) WRO Room may be constructed based on quantity of machines to be supported by it (3 and above) or a portable one if less than 3 units. Mains voltage: 115, 230V Frequency: 50 to 60 Hz Power consumption: Max 2025W at 230 V | 2023-03-01 | 2023-03-01 | OK | IPD
124 | ELECTROENCEPHALOGRAPH STATION - 32 CHANNEL TROLLEY MOUNTED - WITH ACCESSORIES |  | Neurowork | Equipment is mobile on castor Power Supply: 220V, 50/60Hz Power frequency (50/60 Hz) magnetic field IEC 61000-4-8 | 2023-03-01 | 2023-03-01 | OK | IPD
125 | PROJECTOR - MEETING ROOM - CEILING MOUNTED - COMPLETE | [5 empty] | Is this your scope? | IPD
NA | CT 128 SLICE |  | GE Rev Maxima | Rating The system operates on three-phase power that meets the following specifications: • Voltage: 200 to 240 VAC, 380 to 480 VAC • Capacity: 100 kVA • Frequency: 50 or 60 Hz ± 3 Hz • Maximum power demand = 100 kVA @ 0.85 PF at a selected technique of 140 kV, 515 mA. • Average (continuous) power demand at maximum duty cycle = 20 kVA. • Idle power demand (without rotation and X-ray) = 5.0 kVA. | 2023-07-01 | 2023-07-01 | Pending Full Detailed Drawings | IPD
NA | MRI – 1.5T – SUPER CONDUCTIVE – HELIUM COOLED – WITH R/F CAGE AND SET OF COILS |  | GE Signa Explorer | Electrical Power Supply Recommended Configuration: 3-phase Grounded WYE with Neutral and Ground (5 wire system). Alternate Configuration: 3 phase DELTA with Ground (4 wire). Recommend corner Grounded Delta configuration. Frequency: 50 ± 3.0Hz or 60 ± 3.0 Hz (Local Voltage adoption may be required) Voltage: 480/415/400/380/208/200 Vrms Power Consumption Sleep Mode: 6.6 kW Standby (no scan): 9.82 kW Typical consumption per COCIR standards: 14.94 kW Peak Instantaneous power < 5s: 20.4 kW Continuous sustained power: 17.13 kW Minimum room dimensions Magnet Room: 3.53m x 5.7m Control Console Room: 1.52m x 2.13m (with Equip room) 3.334m x 1.9m (no equip room) Electronics Equipment Room: 2.7m x 2.5m Total Area needed excluding chillers: 31.60m2 | 2023-11-01 | 2023-11-01 | Pending Full Detailed Drawings | IPD
NA | DIGITAL MAMMOGRAPHY UNIT WITH SHIELDED CONSOLE |  | GE Senographe Crystal Nova | System Power Supply • Input frequency: 50/60 Hz • Input voltage: single-phase 220-230 VAC High voltage generator • Generator type: high frequency single-phase power supply • Power: 5 kW max • mAs range: 0.1 to 500 mAs (depending on kV) • kV range: 20 to 50 kV, in 0.5 kV steps • Generator protection: software monitoring of generator and tube load • Input frequency: 50/60Hz • Input voltage: single-phase 208 / 220-230 / 240 V • Minimum kVA required: 6.25kVA • Stand-by power: 500W Physical Specification The physical specifications of each component are shown as below. • Components: Gantry (308.6kg), Control Station (with Lead glass 97.2kg) and Generator (91.8kg), and separate Detector (6.3kg) weight tolerance is ±10% Gantry (Main unit): 650*1104*2002mm Control Station: 720*616*1853mm Generator: 360*592*690mm Minimum room layout: (3.2 m x 3.5 m = 11.2 m2) | 2023-04-01 | 2023-04-01 | Pending Full Detailed Drawings | IPD
NA | DIRECT DIGITAL X-RAY RADIOGRAPHY AND FLUOROSCOPY |  | GE Definium XR 575 | Minimum Room area: 25m2 Normal Output Power: 50kW (500mA@100kV, 0.1s) Nominal Input Voltages: 380/400/440/480VAC, Three-phase with or without neutral, normal frequency 50/60Hz. Input Current: 110A (Momentary), 9A (Continuous) | 2023-04-01 | 2023-04-01 | Pending Full Detailed Drawings | IPD
NA | LABORATORY EQUIPMENT (FULL SCOPE) | In IPDF13, IPDF14, IPDF15, IPDF16, IPDF17, IPDF18, IPDF19, and in IVF (IPDG155, IPDG160) |  |  | ??? | ??? | Pending Datasheets Pending Full Detailed Drawings Pending Delivery Schedule | IPD
126 | SERVER UNIT WITH SWITCHES AND CABINET |  |  |  | ??? | ??? | All Details to be Provided | IPD
127 | BED - PEDIATRIC - 2 SECTION - MANUAL - ON CASTORS WITH MATTRESS |  |  |  | ??? | ??? | All Details to be Provided | IPD
128 | DEFIBRILLATOR - UNMONITORED - PEDIATRIC PADDLES |  |  |  | ??? | ??? | All Details to be Provided | IPD
129 | BABY WARMER NICU |  |  |  | ??? | ??? | All Details to be Provided | IPD
130 | VENTILATOR - NEONATAL HFO |  |  |  | ??? | ??? | All Details to be Provided | IPD
131 | INCUBATOR - TRANSPORT - ICU SERVICES - GAS CYLINDER - BATTERY POWERED |  |  |  | ??? | ??? | All Details to be Provided | IPD
132 | INCUBATOR - NICU TR AND REV TR AND VARIABLE HEIGHT ELECTRIC |  |  |  | ??? | ??? | All Details to be Provided | IPD
133 | DEFIBRILLATOR - NEONATAL PADDLES WITH MOINTORING AND PACE |  |  |  | ??? | ??? | All Details to be Provided | IPD
134 | MONITOR - NEONATAL NICU |  |  |  | ??? | ??? | All Details to be Provided | IPD
135 | SINK - CLEANER - SINGLE |  |  |  | ??? | ??? | All Details to be Provided | IPD
136 | INCUBATOR - NICU - ELECTRONIC TEMP CONTROL - FIXED HEIGHT |  |  |  | ??? | ??? | All Details to be Provided | IPD
137 | VENTILATOR - NEONATAL |  |  |  | ??? | ??? | All Details to be Provided | IPD
138 | CPAP MACHINE - NEONATAL |  |  |  | ??? | ??? | All Details to be Provided | IPD
139 | PHOTOTHERAPY UNIT ON CASTORS |  |  |  | ??? | ??? | All Details to be Provided | IPD
140 | INFANT WEIGHING SCALE - ELECTRONIC |  |  |  | ??? | ??? | All Details to be Provided | IPD
141 | RESUSCITATION STATION/BABY WARMER |  |  |  | ??? | ??? | All Details to be Provided | IPD
142 | BED - DELIVERY - HEIGHT ADJUSTABLE - HYDRAULIC - WITH ACCESSORIES |  |  |  | ??? | ??? | All Details to be Provided | IPD
143 | MONITOR - FOETAL TWIN CAPABILITY |  |  |  | ??? | ??? | All Details to be Provided | IPD
144 | BABY COT - TRANSPARENT PRESPEX DESIGN ON CASTORS |  |  |  | ??? | ??? | All Details to be Provided | IPD
145 | RADIANT WARMER - SIMPLE |  |  |  | ??? | ??? | All Details to be Provided | IPD
146 | HEATER - BABY MILK BOTTLES - 6 HOLDER |  |  |  | ??? | ??? | All Details to be Provided | IPD
147 | TABLE - OPERATING - GENERAL SURGERY - ELECTRIC/HYDRAULIC |  |  |  | ??? | ??? | All Details to be Provided | IPD
148 | MONITOR - PATIENT - SURGICAL SUITE - ANAESTHETIC GASES MONITORING |  |  |  | ??? | ??? | All Details to be Provided | IPD
149 | ELECTROSURGICAL UNIT 300 WATTS - SINGLE SURGEON |  |  |  | ??? | ??? | All Details to be Provided | IPD
150 | CRYOSURGERY UNIT - SET OF ACCESSORIES |  |  |  | ??? | ??? | All Details to be Provided | IPD
151 | ENDOSCOPE WASHER - DISINFECTOR |  |  |  | ??? | ??? | All Details to be Provided | IPD
152 | STRETCHER - PROCEDURAL |  |  |  | ??? | ??? | All Details to be Provided | IPD
153 | ELECTROSURGICAL UNIT 120 WATTS - WITH FUME EXTRACTION |  |  |  | ??? | ??? | All Details to be Provided | IPD
154 | FLEXIBLE ENDOSCOPES TOWER WITH CAMERA AND LIGHT SOURCE (VIDEO) |  |  |  | ??? | ??? | All Details to be Provided | IPD