## SOH OPD Tags
SOH OPD MEDICAL EQUIPMENT/FURNITURE TAGGING
Tag No. | Equipment Description | Additional Comments | Brand name | Electrical Specifics | Procurement/ Delivery Schedule | Installation/ Commisioning Schedule | TANIT Comments for A&K
1 | PC DESKTOP 15" MONITOR ALL IN ONE WITH UPS 500 VA | [5 empty] | To be assumed by Electrical Engineer
2 | PRINTER NETWORK BASED LASER B/W | [5 empty] | To be assumed by Electrical Engineer
3 | TV 40" - LED - WALL MOUNTED - WITH SUPPORT | [5 empty] | To be assumed by Electrical Engineer
4 | EXAMINATION HALOGEN LIGHT - ON CASTORS |  |  | Equipment is mobile on wheels, Power supply: 220V, 50Hz, 20W | 2023-01-20 | 2023-02-01 | Supplier only shared Catalogue. Electrical Specifications indicated here
5 | DIAGNOSTIC STATION - WALL MOUNTED - OPHTHALMOSCOPE - OTOSCOPE - NIBP - THERMOMETER |  | Welch Allen | Equipment is wall mounted; Electrical specifications: 100 – 240 V AC, 50 – 60 Hz, 1.5 – 0.8 A | 2023-01-20 | 2023-02-01 | Please check attached Specs for wall mounting height
6 | ULTRA SLIM LIGHT BOX 41X36 CM | [5 empty] | Pending From Supplier
7 | ECG - 12 CHANNEL - ON TROLLEY | [5 empty] | Pending From Supplier
8 | DEFIBRILLATOR - UNMONITORED |  | Progetti | Equipment is mobile, uses rechargeable battery. AC charger power supply, Input: 100 ~ 240V AC 50/60Hz max. Battery pack: 14.4V – 2Ah Ni-Mh battery (internal rechargeable) charging time maximum 3 hours | 2023-01-20 | 2023-02-01 | Please also check attached
9 | TREADMILL ELEVATING WITH PC CONTROL |  | GE Case + T2100 Treadmill | 200 to 240 VAC, 50/60 Hz, single-phase, 20 amperes, NEMA 6-20 R wall socket (dedicated circuit recommended | 2023-01-20 | 2023-02-01 | Please also check attached
10 | ULTRASOUND UNIT - OB/GYN 3/4D VOLUMETRIC WITH DVD RECORDER |  | Voluson S10 | 13A Socket outlet wired with >3 core 2.5mm2 cable, Power Consumption: 900VA Earthing system of less than 2ohms, Voltage 220-240VAC. single phase UPS for Ultrasound; Dedicated AC Power outlet must be within reach of the system without extension cords within 1m of the system. Airconditional Availability | 2023-01-01 | 2023-01-01 | Supplier only shared Catalogue. Electrical Specifications indicated here
11 | PLETHYSMOGRAPH CABINET - COMPACT SIZE WITH CHAIR | [5 empty] | Pending From Supplier
12 | ELECTROENCEPHALOGRAPH STATION - 32 CHANNEL TROLLEY MOUNTED - WITH ACCESSORIES |  | Neurowork | Power frequency (50/60 Hz) magnetic field IEC 61000-4-8 | 2023-01-01 | 2023-02-01 | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
13 | ENT STATION WITH CHAIR AND INSTRUMENTS - VIDEO MONITOR | Total station inculding Chair | Euroclinic | 220/240 V(110/115 VAC) | 2023-01-01 | 2023-02-01 | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
14 | ENT MICROSCOPE |  | Leica | Equipment is mobile on caster wheels. Power supply: 100-240V AC, 50/60Hz. Power consumption: 1200VA | 2023-01-01 | 2023-02-01 | Please also check attached
15 | ELECTROSURGICAL UNIT 200 WATTS - ENT | On ENT Station |  |  |  |  | Pending From Supplier
16 | LAMP - DESKTOP | [5 empty] | To be assumed by Electrical Engineer
17 | WATER DISPENSER | [5 empty] | To be assumed by Electrical Engineer
18 | FRIDGE - SMALL | [5 empty] | To be assumed by Electrical Engineer
18.1 | MICROWAVE SMALL | [5 empty] | To be assumed by Electrical Engineer
18.2 | SCREENER - VISUAL FIELD COMPUTER BASED | May be a TV or a projector depending on the brand |  |  |  |  | Pending From Supplier
19 | CEPHALO SCAN CBCT |  | Belmont | Power Requirement x-ray will operate on a power supply of 120 + 10 VAC. A three wire GROUNDED circuit, separately connected to the central distribution panel with an over current protection device rate for 20 amperes. Recommended wire size is 12 AWG. But if the wire run distance is to exceed 50 feet, 10 AWG is required. For wire run distance in excess of 75 feet, upto 125 feet, 8 AWG is required. Line voltage regulation must exceed 3 % at 12.5 amps. | 0 | 2023-02-01 | Please also check attached
20 | DENTAL CHAIR BUILT IN COMPRESSOR COMPLETE |  | Olsen | Equipment is powered by a Motor: 24V, 46W, 45rpm | 2023-01-01 | 2023-02-01 | Supplier only shared Catalogue. Electrical Specifications indicated here
21 | DENTAL X-RAY UNIT WALL MOUNTED |  | Belmont | Maximum Length Reach of equipment : 84-1/8” with 40.0” arm Power Supply Electrical requirement Power supply of 120VAC ± 10% (108 ~ 132V AC) with a 3 wire GROUNDED circuit, separately connected to the central distribution panel with an over current protection device rated for 15 amperes. Recommended wire size is 14 AWG ; but if the wire run distance is to exceed 50 feet (15m), 12 AWG is required. For wire run distances in excess of 75 feet (23m) 10 AWG is required. Line voltage regulation must be within 0 ~ 3 % at 6.5 amps. Please check drawing on separate sheet | 2023-01-01 | 2023-01-01 | Please also check attached
22 | DENTAL CURING LIGHT - CORDLESS - LED |  | Woodpecker | Equipment is mobile, uses Rechargeable lithium battery. Power input: 100V to 240V, 50/60Hz, 10VA | 2023-01-01 | 2023-02-01 | Please also check attached
23 | AUTOCLAVE - DENTAL INSTRUMENTS 24L |  | Tuttnauer/ 2840EKA D(28L) | Voltage/ Frequency (Hz): 230V/60Hz Power: 1400W 6A Equipment is Table Top with Dimension 630*530*440mm | 2023-01-01 | 2023-02-01 | Supplier only shared Catalogue. Electrical Specifications indicated here
24 | LASER - DENTAL SURGICAL/DIODE/TABLETOP FOR SOFT TISSUE |  | Zolar | Equipment is table-top, uses Battery AC100-240V 50/60Hz Wireless Pedal 1.5V Battery | 2023-01-01 | 2023-02-01 | Please also check attached
25 | AMALGAMATOR |  | Henry Schein | Equipment is Table-top Input Voltage: 110V, 60Hz Power Output: 40W | 2023-01-01 | 2023-02-01 | Please also check attached
26 | RETINAL CAMERA / OCT OPHTHALMOSCOPE / EYE F-ANGIO / TABLETOP |  | Topcon | Equipment is Table-top Input Voltage: 110=240V, 50-60Hz Power Consumption: 400VA (maximum) 100VA (normal) Dimensions W*D*H: (274*508*566mm) | 2023-01-01 | 2023-02-01 | Please also check attached
27 | PROJECTOR - VISUAL ACUITY | [5 empty] | Pending From Supplier
28 | AUTO REFRACTOMETER/KERATOMETER | [5 empty] | Pending From Supplier
28.1 | LENSOMETER - TABLETOP | [5 empty] | Pending From Supplier
29 | OPHTHALMOLOGY CHAIR - COMPLETE (RETINOSCOPE - SLIT LAMP - TONOMETER) |  | Light Box |  |  |  | Pending From Supplier
30 | LED VIDEO COLPOSCOPE - 3.75X/7X/15X > 28000 LUX - ON TROLLEY |  | Ecleris/ C100-FID Colposcope | Equipment is mobile on caster wheels. Power supply: 100-240V AC, 50/60Hz. | 2023-01-01 | 2023-02-01 | Supplier only shared Catalogue. Electrical Specifications indicated here
31 | BOOTH - AUDIOMETRY - COMPACT WITH VIEWPORT AND CHAIR - 1.2X1.2M | [5 empty] | Pending From Supplier
32 | ERGOMETER |  | GE | Power Supply: 100 to 240 V, 50 to 60 Hz Power Consumption: 60VA max | 2023-01-01 | 2023-02-01 | Please also check attached
33 | ELECTROTHERAPY UNIT | [5 empty] | Pending From Supplier
33.1 | PRESSOTHERAPY UNIT | [5 empty] | Pending From Supplier
33.2 | VACCUM THERAPY | [5 empty] | Pending From Supplier
33.3 | INFRARED RADIANT | [5 empty] | Pending From Supplier
33.4 | MAGNETOTHERAPY | [5 empty] | Pending From Supplier
33.5 | CONTINUOUS PASSIVE MOTION UPPER LIMBS |  |  | Power supply: 220V AC, 50Hz. Input Power: <60VA | 2023-01-01 | 2023-02-01 | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
33.6 | CONTINUOUS PASSIVE MOTION LOWER LIMBS | [5 empty] | Pending From Supplier
33.7 | HYDROCOLLECTOR - COLD COMPRESS |  |  | Power Supply: 100 to 240 V, 50 to 60 Hz Power Consumption: 1500W max Dimension: 67*40*84cm | 2023-01-01 | 2023-02-01 | Please also check attached
33.8 | HYDROCOLLECTOR - HOT COMPRESS |  |  |  | 2023-01-01 | 2023-02-01 | Please also check attached
34 | ULTRASOUND UNIT - CARDIAC - WITH ASSESSMENT SOFTWARE - 3 PROBE |  | GE Vivid S7 | 13A Socket outlet wired with >3 core 2.5mm2 cable, Power Consumption: 0.5-1A Earthing system of less than 2ohms, Voltage 240 VAC single phase UPS for Ultrasound Airconditional Availability | 2023-01-01 | 2023-01-01 | Supplier only shared Catalogue. Electrical Specifications indicated here
35 | COMPUTER-BASED URODYNAMIC SYSTEM ON CASTORS | [5 empty] | Pending From Supplier
36 | URODYNAMIC CHAIR | [5 empty] | Pending From Supplier
37 | REFRIGERATOR 500 L |  | Haier | Equipment is mobile on caster wheels. Power supply: 100-240V AC, 50/60Hz. | 2023-01-01 | 2023-01-01 | Please also check attached
38 | PORCELAIN FURNACE WITH VACCUM PUMP | Supplier to provide exact final positioning | European |  |  |  | Pending From Supplier
39 | BURNOUT FURNACE |  | European |  |  |  | Pending From Supplier
40 | CASTING MACHINE - MANUAL |  | European |  |  |  | Pending From Supplier
41 | SANDBLASTER |  | European |  |  |  | Pending From Supplier
42 | MICRO SANDPLASTER |  | European |  |  |  | Pending From Supplier
43 | VIBRATOR |  | European |  |  |  | Pending From Supplier
44 | HANGING MOTOR WITH HANPIECE |  | European |  |  |  | Pending From Supplier
45 | MICRO MOTOR |  | European |  |  |  | Pending From Supplier
46 | VACCUM MIXER |  | China |  |  |  | Pending From Supplier
47 | PINDEX |  | China |  |  |  | Pending From Supplier
48 | MODEL TRIMMER |  | European |  |  |  | Pending From Supplier
49 | STEAM MACHINE |  | China |  |  |  | Pending From Supplier
50 | MOTOR CENTRIFUGE |  | UK |  |  |  | Pending From Supplier
51 | POLISHING MOTOR |  | European |  |  |  | Pending From Supplier
52 | INSTRUMENT CLEANER | [5 empty] | Pending From Supplier
53 | SEALING MACHINE | [5 empty] | Pending From Supplier
54 | ORTHODONTIC SPOT WELDER | [5 empty] | Pending From Supplier
55 | VACUUM FORMER TABLETOP DENTAL LAB | [5 empty] | Pending From Supplier