## SOH OPD Tags
SOH OPD MEDICAL EQUIPMENT/FURNITURE TAGGING
Tag No. | Equipment Description | Supply Source (UPS / Normal) | Additional Comments | Brand name | Electrical Specifics | TANIT Comments for A&K
1 | PC DESKTOP 15" MONITOR ALL IN ONE WITH UPS 500 VA | UPS |  |  |  | To be assumed by Electrical Engineer
2 | PRINTER NETWORK BASED LASER B/W | UPS (Duplex Socket With Station) |  |  |  | To be assumed by Electrical Engineer
3 | TV 40" - LED - WALL MOUNTED - WITH SUPPORT | Normal (TV Reception Area) |  |  |  | To be assumed by Electrical Engineer
4 | EXAMINATION HALOGEN LIGHT - ON CASTORS | UPS |  |  | Equipment is mobile on wheels, Power supply: 220V, 50Hz, 20W | Supplier only shared Catalogue. Electrical Specifications indicated here
5 | DIAGNOSTIC STATION - WALL MOUNTED - OPHTHALMOSCOPE - OTOSCOPE - NIBP - THERMOMETER | UPS |  | Welch Allen | Equipment is wall mounted; Electrical specifications: 100 – 240 V AC, 50 – 60 Hz, 1.5 – 0.8 A | Please check attached Specs for wall mounting height
6 | ULTRA SLIM LIGHT BOX 41X36 CM | Normal |  | GF TENDERCARE | Equipment is Table top. Input Power Supply: 100-230V, 50-60Hz, 120VA | Please also check attached
7 | ECG - 12 CHANNEL - ON TROLLEY | UPS |  | GE Mac 2000 | Equipment is mounted on cart to a height of 200mm. Equipment uses AC/DC or Battery Operation. Input Power Supply: 100-240VAC, 47-63Hz, maximum 1.5A. Equipment is also LAN enabled. | Please also check attached
8 | DEFIBRILLATOR - UNMONITORED | UPS |  | Progetti | Equipment is mobile, uses rechargeable battery. AC charger power supply, Input: 100 ~ 240V AC 50/60Hz max. Battery pack: 14.4V – 2Ah Ni-Mh battery (internal rechargeable) charging time maximum 3 hours | Please also check attached
9 | TREADMILL ELEVATING WITH PC CONTROL | UPS |  | GE Case + T2100 Treadmill | 200 to 240 VAC, 50/60 Hz, single-phase, 20 amperes, NEMA 6-20 R wall socket (dedicated circuit recommended | Please also check attached
10 | ULTRASOUND UNIT - OB/GYN 3/4D VOLUMETRIC WITH DVD RECORDER | UPS |  | Voluson S10 | 13A Socket outlet wired with >3 core 2.5mm2 cable, Power Consumption: 900VA Earthing system of less than 2ohms, Voltage 220-240VAC. single phase UPS for Ultrasound; Dedicated AC Power outlet must be within reach of the system without extension cords within 1m of the system. Airconditional Availability | Supplier only shared Catalogue. Electrical Specifications indicated here
11 | PLETHYSMOGRAPH CABINET - COMPACT SIZE WITH CHAIR | UPS |  | nSpire | Input Power Supply: 100-240V ± 10%, 50/60 Hz. Box Dimension: 93*85*172cm Equipment requires external PC connection. | Please also check attached
12 | ELECTROENCEPHALOGRAPH STATION - 32 CHANNEL TROLLEY MOUNTED - WITH ACCESSORIES | UPS |  | Neurowork | Power frequency (50/60 Hz) magnetic field IEC 61000-4-8 | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
13 | ENT STATION WITH CHAIR AND INSTRUMENTS - VIDEO MONITOR | UPS | Total station inculding Chair | Euroclinic | 220/240 V(110/115 VAC) | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
14 | ENT MICROSCOPE | UPS (DUPLEX WITH 15) |  | Leica | Equipment is mobile on caster wheels. Power supply: 100-240V AC, 50/60Hz. Power consumption: 1200VA | Please also check attached
15 | ELECTROSURGICAL UNIT 200 WATTS - ENT | UPS (DUPLEX WITH 14) | On ENT Station | Inspital | Input Power Voltage: 100-240VAC, 50-60Hz Input Power: 750VA Maximum Output Power 200W | Please also check attached
16 | LAMP - DESKTOP | Normal |  |  |  | To be assumed by Electrical Engineer
17 | WATER DISPENSER | Normal |  |  |  | To be assumed by Electrical Engineer
18 | FRIDGE - SMALL | Normal |  |  |  | To be assumed by Electrical Engineer
18.1 | MICROWAVE SMALL | Normal |  |  |  | To be assumed by Electrical Engineer
18.2 | SCREENER - VISUAL FIELD COMPUTER BASED | UPS |  | LIGHT BOX | Equipment is Table top, LAN Connectivity enabled. Input Power Supply: 85 – 263 V AC, 50/60 Hz, 60 VA | Please also check attached
19 | CEPHALO SCAN CBCT | UPS |  | Belmont | Power Requirement x-ray will operate on a power supply of 120 + 10 VAC. A three wire GROUNDED circuit, separately connected to the central distribution panel with an over current protection device rate for 20 amperes. Recommended wire size is 12 AWG. But if the wire run distance is to exceed 50 feet, 10 AWG is required. For wire run distance in excess of 75 feet, upto 125 feet, 8 AWG is required. Line voltage regulation must exceed 3 % at 12.5 amps. | Please also check attached
20 | DENTAL CHAIR BUILT IN COMPRESSOR COMPLETE | UPS (WALL MOUNTED NEAR THE CHAIR) - IF NEEDED WE CAN CHANGE TO FLOOR MOUNTED STUB UP) |  | Olsen | Equipment is powered by a Motor: 24V, 46W, 45rpm | Supplier only shared Catalogue. Electrical Specifications indicated here
21 | DENTAL X-RAY UNIT WALL MOUNTED | UPS (Duplex With Item 20) |  | Belmont | Maximum Length Reach of equipment : 84-1/8” with 40.0” arm Power Supply Electrical requirement Power supply of 120VAC ± 10% (108 ~ 132V AC) with a 3 wire GROUNDED circuit, separately connected to the central distribution panel with an over current protection device rated for 15 amperes. Recommended wire size is 14 AWG ; but if the wire run distance is to exceed 50 feet (15m), 12 AWG is required. For wire run distances in excess of 75 feet (23m) 10 AWG is required. Line voltage regulation must be within 0 ~ 3 % at 6.5 amps. Please check drawing on separate sheet | Please also check attached
22 | DENTAL CURING LIGHT - CORDLESS - LED | UPS (With 24 & 25) DUPLEX |  | Woodpecker | Equipment is mobile, uses Rechargeable lithium battery. Power input: 100V to 240V, 50/60Hz, 10VA | Please also check attached
23 | AUTOCLAVE - DENTAL INSTRUMENTS 24L | Normal |  | Tuttnauer/ 2840EKA D(28L) | Voltage/ Frequency (Hz): 230V/60Hz Power: 1400W 6A Equipment is Table Top with Dimension 630*530*440mm | Supplier only shared Catalogue. Electrical Specifications indicated here
24 | LASER - DENTAL SURGICAL/DIODE/TABLETOP FOR SOFT TISSUE | UPS (With 24 & 25) DUPLEX |  | Zolar | Equipment is table-top, uses Battery AC100-240V 50/60Hz Wireless Pedal 1.5V Battery | Please also check attached
25 | AMALGAMATOR | UPS (With 24 & 25) DUPLEX |  | Henry Schein | Equipment is Table-top Input Voltage: 110V, 60Hz Power Output: 40W | Please also check attached
26 | RETINAL CAMERA / OCT OPHTHALMOSCOPE / EYE F-ANGIO / TABLETOP | UPS |  | Topcon | Equipment is Table-top Input Voltage: 110=240V, 50-60Hz Power Consumption: 400VA (maximum) 100VA (normal) Dimensions W*D*H: (274*508*566mm) | Please also check attached
27 | PROJECTOR - VISUAL ACUITY | UPS (DUPLEX WITH WORKSTATION) |  | LIGHT BOX AUTOMATED ACUITY CHAT | Equipment is Table top, can also be mounted on wall. Input Power Supply: 115-230 VAC 50/60 Hz | Please also check attached
28 | AUTO REFRACTOMETER/KERATOMETER | UPS (ADDED IN NEW REVISED LAYOUT) |  | REICHART | Equipment is Table top. Input Power Supply: 100-240V, 50/60Hz, 60VA | Please also check attached
28.1 | LENSOMETER - TABLETOP | UPS (DUPLEX WITH WORKSTATION) |  | LIGHT BOX | Equipment is Table top. Input Power Supply: 6VAC Adapter 220V | Please also check attached
29 | OPHTHALMOLOGY CHAIR - COMPLETE (RETINOSCOPE - SLIT LAMP - TONOMETER) | UPS |  | Light Box | Input Power Supply: 110-230V, 50/60Hz | Please also check attached
30 | LED VIDEO COLPOSCOPE - 3.75X/7X/15X > 28000 LUX - ON TROLLEY | Normal | PhysioTherapy Zone, No need for UPS (Not essential and not critical load) | Ecleris/ C100-FID Colposcope | Equipment is mobile on caster wheels. Power supply: 100-240V AC, 50/60Hz. | Supplier only shared Catalogue. Electrical Specifications indicated here
31 | BOOTH - AUDIOMETRY - COMPACT WITH VIEWPORT AND CHAIR - 1.2X1.2M | Normal |  | sibelmed | Input Power Supply: 220 V 50-60 Hz, 50W | Please also check attached
32 | ERGOMETER | Normal |  | GE | Power Supply: 100 to 240 V, 50 to 60 Hz Power Consumption: 60VA max | Please also check attached
33 | ELECTROTHERAPY UNIT | Normal |  | Doctor life | Input Power Supply: 220 V 50-60 Hz, 50W | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
33.1 | PRESSOTHERAPY UNIT | Normal |  | EME Physio | Input Power Supply: 220 V 50-60 Hz, 50W | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
33.2 | VACCUM THERAPY | Normal |  | Motech | Equipment is mobile, hand-held Input Power Supply: 200 V, 50Hz | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
33.3 | INFRARED RADIANT | Normal |  | Motech | Equipment is mobile, hand-held Input Power Supply: 200 V, 50Hz | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
33.4 | MAGNETOTHERAPY | Normal |  | Chinesport | Equipment is Table-top Input Power Supply: 230 VAC, 50-60 Hz | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
33.5 | CONTINUOUS PASSIVE MOTION UPPER LIMBS | Normal |  |  | Power supply: 220V AC, 50Hz. Input Power: <60VA | No datasheet/Catalogue shared but electrical specifications indicated here. Please let us know if this is sufficient
33.6 | CONTINUOUS PASSIVE MOTION LOWER LIMBS | Normal |  | EME Physio | Input Power Supply: 220V, 50Hz | Please also check attached
33.7 | HYDROCOLLECTOR - COLD COMPRESS | Normal |  |  | Power Supply: 100 to 240 V, 50 to 60 Hz Power Consumption: 1500W max Dimension: 67*40*84cm | Please also check attached
33.8 | HYDROCOLLECTOR - HOT COMPRESS | Normal |  |  |  | Please also check attached
34 | ULTRASOUND UNIT - CARDIAC - WITH ASSESSMENT SOFTWARE - 3 PROBE | UPS |  | GE Vivid S7 | 13A Socket outlet wired with >3 core 2.5mm2 cable, Power Consumption: 0.5-1A Earthing system of less than 2ohms, Voltage 240 VAC single phase UPS for Ultrasound Airconditional Availability | Supplier only shared Catalogue. Electrical Specifications indicated here
35 | COMPUTER-BASED URODYNAMIC SYSTEM ON CASTORS | UPS (DUPLEX WITH 36) |  | creomidical | Equipment is moveable on castor. LAN Connection available. Input Power Supply: 220V, 50Hz | Supplier only shared Catalogue. Electrical Specifications indicated here
36 | URODYNAMIC CHAIR | UPS (DUPLEX WITH 35) |  | Creo Medical | Input Power Supply: 220-240V, 50/60Hz | Please also check attached
37 | REFRIGERATOR 500 L | UPS |  | Haier | Equipment is mobile on caster wheels. Power supply: 100-240V AC, 50/60Hz. | Please also check attached
38 | PORCELAIN FURNACE WITH VACCUM PUMP | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 39 | Supplier to provide exact final positioning | Phoenix 7 | Input Power Supply: 220V, 1.6W | Please also check attached
39 | BURNOUT FURNACE | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 38 |  | UMG | Input Power Supply: 240V, 13A, 3.1kW, 1 Ph, 50/60Hz; Power Consumption: 4000W | Please also check attached
40 | CASTING MACHINE - MANUAL |  |  | Mancast | Equipment is table top, non electrical. | Please also check attached
41 | SANDBLASTER | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 41, 42 & 43 |  | Aixin | Input Power Supply: 220V, 50/60Hz Power Consumption: <8W | Please also check attached
42 | MICRO SANDPLASTER | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 41, 42 & 43 |  | Optiblast | Equipment is table top, mobile. Input Power Supply: 220V, 50Hz | Please also check attached
43 | VIBRATOR | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 41, 42 & 43 |  | European | Equipment is table top Input Power Supply: 220V, 50Hz | Please also check attached
44 | HANGING MOTOR WITH HANPIECE | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 44, 45 & 47 |  | European | Equipment is mobile with rechargeable battery 40V, 2.5Ah Input Power Supply: 220V, | Please also check attached
45 | MICRO MOTOR | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 44, 45 & 47 |  | European | Equipment is attached to motor | Please also check attached
46 | VACCUM MIXER | UPS |  | China | Equipment is table top Input Power Supply: 110/230V, 50/60Hz. Motor: 320RPM | Please also check attached
47 | PINDEX | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 44, 45 & 47 |  | China | Equipment is table top Input Power Supply: 230V, 50Hz, 2500rpms Nominal Power: 100W | Please also check attached
48 | MODEL TRIMMER | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 48 & 49 |  | European | Power Supply: 220V/110V 50/60Hz Power Consumption: 370W Rotational Speed: 2,800RPM | Please also check attached
49 | STEAM MACHINE | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 48 & 49 |  | China | Power Supply: 120/220V Power Consumption: 10amps | Please also check attached
50 | MOTOR CENTRIFUGE | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 50, 51 & 52 |  | UK | Power Supply: 220V | Please also check attached
51 | POLISHING MOTOR | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 50, 51 & 52 |  | European | Power Supply: 220V, 50/60Hz | Please also check attached
52 | INSTRUMENT CLEANER | UPS (ADDED IN NEW REVISED LAYOUT) - Duplex With 50, 51 & 52 |  | MULIT ENZYMATIC | Consumable, Non-Electrical Item | Consumable, Non-Electrical Item
53 | SEALING MACHINE | UPS (ADDED IN NEW REVISED LAYOUT) |  | QASA | Equipment is Table top Power Input: max 420W | Please also check attached
54 | ORTHODONTIC SPOT WELDER | UPS (ADDED IN NEW REVISED LAYOUT) |  | Vimel | Equipment is Table top Power Input Voltage: 220V, 50/60Hz Supply current on load: 15A Power: 1000W | Please also check attached
55 | VACUUM FORMER TABLETOP DENTAL LAB | UPS |  | UMG | Equipment is Table top Power Input Voltage: 220V, 50/60Hz Power: 350W | Please also check attached